## Worksheet 1
Variable |  | Required | Desired
Salinity |  | AutoSal | (Sub-) standard used regularly
Temperature constant
Glass bottles
Dissolved oxygen |  | Winkler | Draw temperature used for mass calculation if difference to in situ temperature > 2.5 °C
Titration reagent assessed using CSK/OSIL primary standard
Nutrients | All except silicate | Autoanalyzer If stored: frozen upright | Carrier solution documented Calibrated against reference material
| Silicate | Autoanalyzer If stored and concentrations are above 40 µmol L−1: | Carrier solution documented
|  | Poisoned and kept cold | Calibrated against reference material
Dissolved inorganic carbon |  | Coulometry | Calibrated against certified reference material (Andrew Dickson, SIO)
If stored: poisoned, kept in dark and cool location
|  | Potentiometric (closed cell) Calibrated against certified reference material (Andrew Dickson, SIO); If stored: poisoned and kept in a cold, dark location* | Not applicable
Total alkalinity |  | Potentiometric titration (multi-step) | Open-cell or curve-fitting method documented
Calibrated against certified reference material (Andrew Dickson, SIO)
If stored: poisoned, kept in dark and cool location
|  | Spectrophotometric | Indicator dye documented
Calibrated against certified reference material (Andrew Dickson, SIO)
If stored: poisoned, kept in dark and cool location
pH |  | Spectrophotometric with scale and temperature reported | Indicator dye: m-cresol purple
Indicator dye: purified
If dye is not purified: correction applied to impurities
Partial pressure of CO2 |  | Gas chromatography | Temperature and standard reported
If stored: poisoned, kept in dark and cool location
|  | Infrared-based system | Temperature and standard reported
If stored: poisoned, kept in dark and cool location
Particulate matter | Carbon and nitrogen | High-temperature combustion with reported filter volume and pore size | Dried filters (60 °C)
Standards reported
| Phosphorus | Ash hydrolysis with reported filter volume and pore size | Dried filters (60 °C)
Standards reported
Dissolved organic carbon |  | High-temperature combustion | Filtered
If stored: frozen or acidified and refrigerated
Calibrated against reference material (Dennis Hansell, University of Miami)